## notifications Output
Group and Code | Group | Description | Recommended Action
5:1 | PCIe Driver | The Driver encountered an unexpected Windows status. This is a fatal software error requiring a system restart. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:2 | PCIe Driver | The Gemini PCI Express Interface Card failed its hardware memory POST. This is a fatal hardware error indicating a failed Gemini PCI Express card. | Hardware failure. Replace the PCI card.
5:3 | PCIe Driver | The Driver encountered an internal failure within the CANbus Heartbeat thread and has terminated this thread. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:4 | PCIe Driver | The Driver encountered an internal failure within the Async notification thread and has terminated this thread. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:5 | PCIe Driver | The Driver encountered an internal failure within the hardware monitoring thread and has terminated this thread. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:6 | PCIe Driver | The Driver global CANbus message pool is empty. This may indicate an internal driver or application problem. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:7 | PCIe Driver | The CANbus hardware encountered a BUS_OFF condition, and the CANbus has been disabled. | Check CANbus connection from AWS to gantry
5:8 | PCIe Driver | The CANbus hardware encountered a transmit hardware overflow condition. This is a fatal software error requiring a system restart. | Check CANbus connection from AWS to gantry
5:9 | PCIe Driver | The CANbus hardware encountered a receive hardware overflow condition. CANbus data was lost. | Check CANbus connection from AWS to gantry
5:10 | PCIe Driver | There was a problem with DMA on the card. This may indicate a failing Gemini PCI Express card or a failing motherboard. | Hardware failure. Call TS.
5:11 | PCIe Driver | There was a Link Fault. This may indicate a cabling problem, a failing Gemini PCI Express card or a failing motherboard. | Check cabling between AWS and gantry. If cable is ok, use loopback to determine which end of fiber is defective.
5:12 | PCIe Driver | There was a SerialLite Protocol Error. This is most likely due to a PCI card firmware bug or a readout sequence bug. | A software error has occurred. Check that the correct PCI firmware was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:13 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. ( DrvnMajPciMemPaneIndex ) | Hardware failure. Replace the PCI card.
5:14 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. ( DrvnMajPciMemChunkOverrange ) | Hardware failure. Replace the PCI card.
5:15 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. ( DrvnMajPciMemChunkCount ) | Hardware failure. Replace the PCI card.
5:16 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. ( DrvnMajPciMemExposeOverwrite ) | Hardware failure. Replace the PCI card.
5:17 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. | Hardware failure. Replace the PCI card.
5:18 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. | Hardware failure. Replace the PCI card.
5:19 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. | Hardware failure. Replace the PCI card.
5:20 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. | Hardware failure. Replace the PCI card.
5:21 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. | Hardware failure. Replace the PCI card.
5:22 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. ( DrvnMajPciMemWriteHeader ) | Hardware failure. Replace the PCI card.
5:23 | PCIe Driver | PCIe image memory failure. This error may indicate problems with the readout sequence or failing PCI Express card. | Hardware failure. Replace the PCI card.
5:24 | PCIe Driver | Image path failure. This may indicate a failed PCIe Card, possible problem with the Detector, or potential problems with the fiber cable. | Hardware failure. Replace the PCI card.
5:25 | PCIe Driver | There was a problem with the power on the card. This may indicate a failing Gemini PCI Express card. | Hardware failure. Replace the PCI card.
5:26 | PCIe Driver | There was a problem with the power on the card. This may indicate a failing Gemini PCI Express card. | Hardware failure. Replace the PCI card.
5:27 | PCIe Driver | Receive Carrier failed. This most likely is due to the Detector being powered off unexpectedly. This may also indicate possible cabling problems. | Check detector power, PCI configuration, and fiber cabling.
5:28 | PCIe Driver | Image data may be corrupt. This indicates a failed PCIe Card, possible problem with the Detector, or potential problems with the fiber cable. | Check detector power, PCI configuration, and fiber cabling.
5:29 | PCIe Driver | There was a problem with image memory on the card. This may indicate a failing Gemini PCI Express card or problems with the readout sequence. | Hardware failure. Replace the PCI card.
5:30 | PCIe Driver | There was a problem with image memory on the card. This may indicate a failing Gemini PCI Express card or problems with the readout sequence. | Hardware failure. Replace the PCI card.
5:31 | PCIe Driver | There was a problem with image memory on the card. This may indicate a failing Gemini PCI Express card. | Hardware failure. Replace the PCI card.
5:32 | PCIe Driver | There was a problem with image memory on the card. This may indicate a failing Gemini PCI Express card or problems with the readout sequence. | Hardware failure. Replace the PCI card.
5:33 | PCIe Driver | The CANbus hardware is encountering unexpected CANbus errors. | Check CANbus connection. Call TS if problem persists.
5:34 | PCIe Driver | The Driver is incompatible with the version of PCI Express Firmware currently loaded. Please update the firmware with a compatible version. | Insure that the correct version of firmware and driver are loaded.
5:35 | PCIe Driver | There was a SerialLite Protocol Error. The address is not the start event, end event or image address. | Call TS
5:36 | PCIe Driver | There was a SerialLite Protocol Error. The start event address was received without having both SOP and EOP set. | Call TS
5:37 | PCIe Driver | There was a SerialLite Protocol Error. The end event address was received without having both SOP and EOP set. | Call TS
5:38 | PCIe Driver | There was a SerialLite Protocol Error. The image address was received having both SOP and EOP set. | Call TS
5:39 | PCIe Driver | There was a SerialLite Protocol Error. An incorrect combination of SOP, EOP, and address has occurred. | Call TS
5:40 | PCIe Driver | There was a SerialLite Protocol Error. The SerialLite link has failed. | Call TS
5:41 | PCIe Driver | There was a SerialLite Protocol Error. An unknown error condition has occurred or more than one of SerialLite errors has occurred. | Call TS
5:42 | PCIe Driver | An Extended Announcement packet was detected with a zero packet ID. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:43 | PCIe Driver | An Extended Announcement packet was detected with a zero packet count. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:44 | PCIe Driver | An Extended Announcement packet was detected with a packet ID greater than the total packet count. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:45 | PCIe Driver | An Extended Announcement packet was detected that was not the last packet, but had less than 8 bytes of data. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:46 | PCIe Driver | An Extended Announcement message was received with duplicate packet IDs. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:47 | PCIe Driver | An ACCEPT message was received that was malformed. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:48 | PCIe Driver | A DENY message was received that was malformed. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:49 | PCIe Driver | A message was received with a NULL Destination ID. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:50 | PCIe Driver | A message was received with a NULL Source ID. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:51 | PCIe Driver | A message was received with an invalid Primary Type. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
5:52 | PCIe Driver | Problems with the Image link on the card. This may be due to the Detector being powered off unexpectedly. This may also indicate cabling problems. | Check detector power, PCI configuration, and fiber cabling.
5:53 | PCIe Driver | An Extended Announcement message was received with some packets containing invalid Packet Totals. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
7:1 | M35 Subsys | The imaging subsystem failed to initialize. | Check for M35 initialization errors and that all of the Detector configuration files are present in the M35 Directory. Check the logs for the AWS passing a null buffer, and the M35 image buffer not being ready. Check that the image index is properly configured and valid. Check the logs for Exposure Index, AEC Sensor and ROI region errors.
7:2 | M35 Subsys | An asynchronous error was received from the imaging subsystem. | The PCIe driver cannot be opened. Check for PCIe errors, and the PCIe driver configuration. Check the fiber cable and connectors.
7:3 | M35 Subsys | An error was returned by the imaging subsytem when reading the image data | Check that the six detector defect files are present and valid in the M35_DET_DIR.
7:4 | M35 Subsys | The imaging subsystem returned an error when trying to ARM | Check to see if the PCIe card was already armed, and for other PCIe card issues. Check the fiber cable and connectors.
7:5 | M35 Subsys | The imaging subsystem returned an error when trying to DISARM. | Check the logs for PCIe errors and status. Check that the PCIe driver is working correctly. Check the fiber cable and connectors.
7:6 | M35 Subsys | The imaging subsystem returned an error from the reload files request | Check the logs to see if the M35 is already initialized. Check the parameters passed from the AWS to the M35 for validity. Check that the M35_DIR and M35_DET_DIR are present and populated. Check the operating system for missing M35 dll files. Check the log files to for a M35 checksum failure.
7:7 | M35 Subsys | The imaging subsystem returned an error while configuring the detector. | Check logs to see if the M35 has initialized, or the M35 is currently capturing an image. Check to make sure the M35 has correctly loaded the defect map and gain files, and has the proper Tomo configuration. Finally check for the AWS passing invalid parameters to the M35.
7:8 | M35 Subsys | The imaging channel verification that the M35 performs directed by the AWS at startup has failed. | Hardware failure. Call TS.
8:1 | Detector | TFS component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:2 | Detector | (obsolete) ZLib component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:3 | Detector | Shell component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:4 | Detector | Task component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:5 | Detector | Logger component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:6 | Detector | CAN I/O component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:7 | Detector | Sysinit component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:8 | Detector | AEC component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:9 | Detector | ROS component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:10 | Detector | XSM component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:11 | Detector | Array component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:12 | Detector | An error has occurred in the DSPI communications channel to the Analog PCB, or Simultaneous QSPI and DSPI operations were attempted, forcing a sharing violation. | Check that the correct firmware is installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:13 | Detector | Jobs component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:14 | Detector | (obsolete) Jbi component requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:15 | Detector | (obsolete) Application requires attention | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:16 | Detector | DSS Boot Loader failed to load the application software. | Check that the correct Detector software was installed successfully. Otherwise escalate. Logfile Textgram includes details.
8:17 | Detector | (obsolete) DSS Application Loader failed to load the application...software. | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:18 | Detector | This is issued whenever any image parameters which require a reboot are changed. | Reboot the detector.
8:19 | Detector | Issued during POST if initialization fails. This can be either a software or a hardware problem. Logfile will include more details about where during POST Detector failed. | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:20 | Detector | Issued if a software module fails to terminate correctly during the DET's shutdown sequence. | Nothing. This fault shouldn't matter since the DET is rebooting anyway.
8:21 | Detector | (obsolete) TFS component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:22 | Detector | (obsolete) ZLib component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:23 | Detector | Shell component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:24 | Detector | (obsolete) Task component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:25 | Detector | Logger component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:26 | Detector | CAN I/O component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:27 | Detector | (obsolete)Sysinit component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:28 | Detector | AEC component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:29 | Detector | ROS component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:30 | Detector | XSM component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:31 | Detector | A software error occurred during array initialization. | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:32 | Detector | Jobs component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:33 | Detector | (obsolete) JBI component fault | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:34 | Detector | (obsolete) Application fault | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
8:35 | Detector | (obsolete) Invalid ROS instruction | Check that the correct readout sequence was installed. Otherwise escalate to TS.
8:36 | Detector | Detector analog PCB selftest error. Differential communication PLL is not locked on analog PCB. | Make sure analog board connected and powered. Replace analog PCB.
8:37 | Detector | FPGA selftest failed. | Check that the correct FPGA firmware is loaded. If unresolved, replace CPU board.
8:38 | Detector | EL panel self-test failed. See textgram in logfile for details. | Replace Detector.
8:39 | Detector | HV supply selftest or calibration error has occurred. | Check HV Power Supply. If unresolved, escalate.
8:40 | Detector | An error occurred when trying to access the EEPROM. See textgram in logfile for details. | Check cable between CPU board and analog board. If unresolved, escalate.
8:41 | Detector | An error occurred reading the EL sensing circuit. | Replace CPU board (DSS) or MAC board (HDT).
8:42 | Detector | SPI communication timed out. Detector unable to communicate to internal components. | Hardware failure. DSS - Possible bad CPU or analog board; HDT - Possible bad MAC board
8:43 | Detector | Gatewalk timed out. Detector unable to complete gate clocking. | Replace CPU or analog board (DSS) or MAC board (HDT). If unresolved, escalate.
8:44 | Detector | Read timed out. Detector image readout failed to complete. | Replace CPU or analog board (DSS) or MAC board (HDT). If unresolved, escalate.
8:45 | Detector | An internal hardware communication error occurred. See textgram in logfile for details. | Check cable between CPU board and analog board. If unresolved, replace CPU board (DSS). Replace MAC board (HDT).
8:46 | Detector | A channel-out-of-range error occurred. See textgram in logfile for details. | Check that the correct software was installed using correct instructions. If unresolved, escalate.
8:47 | Detector | An error occurred ramping the HV - the high-voltage power supply is not calibrated. | Check for and address POST failures. If unresolved, escalate.
8:48 | Detector | An error occurred during Detector Power-On Self Test. See textgrams in logfile for specifics of POST failure. | Resolve issue(s) identified in the POST failure log messages.
8:49 | Detector | Detector could not load the ROS file because of an ROS input linkage error. | A software error has occurred. Check that the correct software and ROS were installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:50 | Detector | Detector could not load the state machine file because of a XSM input linkage error. | A software error has occurred. Check that the correct software and ROS were installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:51 | Detector | Detector could not load the ROS because of an ROS binding error. | A software error has occurred. Check that the correct software and ROS were installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:52 | Detector | Detector could not load the state machine file because of a XSM input binding error. | A software error has occurred. Check that the correct software and ROS were installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:53 | Detector | Detector could not link to the XSM because another timer was already linked. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:54 | Detector | Detector could not create the XSM timer because the timer already exists or too many timers have been created. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:55 | Detector | The software has attempted to transmit an extended CAN message with too much data. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:56 | Detector | Requested operation is not allowed when sequencer is enabled | User must stop the readout sequencer before attempting this operation.
8:57 | Detector | An operation was attempted which is not allowed in the current exposure state. | The user must change to the correct exposure state, calibration state or invalid calibration state before attempting this operation.
8:58 | Detector | An invalid calibration index was selected. The AWS or service tool passed an invalid index parameter as part of the can command CAN_DET_AEC_CAL_CLEAR. | Check that the correct Detector software was installed, and that the service software is compatible with this version of detector software. Otherwise escalate. Logfile Textgram includes details.
8:59 | Detector | Detector software failed to open or write either the aec.conf or darkoffset.conf file. | Check to see if the aec.conf and darkoffset.conf files are installed on Detector; path is /tmp/*.conf. Otherwise escalate. Logfile Textgram includes details.
8:60 | Detector | (obsolete) Application setup read failed. | A software error has occurred. Check that the correct software was installed using correct instructions. Otherwise escalate to TS. Logfile Textgram includes details.
8:61 | Detector | An operation was attempted which is not allowed in the current operating mode. | User must change to the correct operational mode before attempting this operation.
8:62 | Detector | An error occurred calculating AEC. See textgram in logfile for details. | Check AEC calibrations.
8:63 | Detector | Issued when Detector temperature is above its high temperature shutdown limit | Check room temperature, DTC operation, Detector cover-fan operation, and air vents. If unresovled, escalate.
8:64 | Detector | Issued when Detector reaches the high temperature warning level. | Check room temperature, DTC operation, Detector cover-fan operation, and air vents. If unresovled, escalate.
8:65 | Detector | The DTC is not able to keep the detector temperature within its setpoint tolerance. | The current Detector temperature exceeds target threshold.
8:66 | Detector | Issued when Detector does not receive current temp from DTC | Check if DTC is running. Check connections to DTC. Ignore if uploading DTC software. If unresovled, escalate.
8:67 | Detector | A software error occurred initializing the temperature limits - values are out of range. | Check for incorrect software version. Check if temperature limits are correctly set. If unresovled, escalate.
8:68 | Detector | ROS attempted an invalid operation. | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:69 | Detector | An array gate-loop or source end-of-scan error occurred. See textgram in logfile for details. | Check that all six flex connections from the hybrids to the Analog PCB are seated and locked. If unresolved, replace the Detector.
8:70 | Detector | An error occurred transmitting readout data from the Analog board to the CPU board. Check logfile textgram for details. | Check that the flat ribbon cable connecting the Analog PCB to the CPU PCB is seated and screwed down firmly. If unresolved, replace the Detector.
8:71 | Detector | An error occurred during transmission over fiber optic link. See textgram in logfile for details. | Check fiber cable between PCIe card and gantry.
8:72 | Detector | An over-voltage error occurred in the high-voltage power supply. | Check that the correct readout sequence is loaded. Check for faulty HV power supply or HV control. Otherwise escalate.
8:73 | Detector | An error occurred ramping the high-voltage power supply. | Check that the correct readout sequence is loaded. Check for faulty HV power supply or HV control. Check textgram in logfile for more details. Otherwise escalate.
8:74 | Detector | Gate driver current is outside of allowed tolerance. | Replace Detector if persistent.
8:75 | Detector | An error occurred accessing lifetime statistics information. | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:76 | Detector | An error occurred reading the Detector identifier - serial number and/or model information in /opt/indentifier.conf is missing or invalid. | Reload /opt/identifier.conf. If unresolved, escalate.
8:77 | Detector | A software error occurred dumping the HV. | Check that software versions are properly installed. If unresolved, escalate.
8:78 | Detector | EL panel run-time verification failed. See textgram in logfile for details. | Check tail connector in pass-through board and connector on CPU board (DSS). Check connector on MAC board. If unresolved, replace Detector (HDT).
8:79 | Detector | An error occurred loading the Detector configuration. | Check that the correct software was installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:80 | Detector | An error occurred loading the AEC sensor table (/etc/sensortab) | Check for incorrect software version or configuration files. Check validity of custom sensortab files (/opt/sensortab). Otherwise escalate.
8:81 | Detector | The goaltab and/or dark-offset settings are invalid for the currently selected modality index. | Perform AEC calibration.
8:82 | Detector | Issued at AEC calculation if the ROS has not verified the selected AEC Cal index | Check for incorrect software version or configuration files. Otherwise escalate. Update ROS to verify AEC index before use.
8:83 | Detector | Invalid AEC sensor region requested. | Check for incorrect software version or configuration files. Otherwise escalate.
8:84 | Detector | An AEC memory error occurred during self-test or sensor readout. | Replace CPU board (DSS) or MAC board (HDT). If unresolved, escalate.
8:85 | Detector | An error occurred during HSS loopback self-test. See textgram in logfile for details. | Escalate as possible hardware error.
8:86 | Detector | The Detector requires rebooting after diagnostics have been performed since the hardware is now in an unspecified state. | Reboot Detector.
8:87 | Detector | (obsolete) Detector has not reached a stable state after startup. | Please wait for Detector to warm up.
8:88 | Detector | The requested operation is not allowed in the current diagnostic mode | The user must change to the correct diagnostic mode before attempting this operation.
8:89 | Detector | An error occurred mounting the Detector filesystem. | Check for incorrect software version or configuration files. Otherwise escalate as a possible hardware error.
8:90 | Detector | An error occurred verifying the Detector filesystem. | Escalate as possible hardware error.
8:91 | Detector | ROS execution was aborted due to run-time error. Typically this points to a bug in either Detector software or ROS program. | Confirm that the DSS and ROS software versions are correct. Otherwise escalate.
8:92 | Detector | ROS execution was self-terminated. | This is not an error by itself, but a normal ROS should never terminate. Check for other error conditions, confirm Detector and ROS software versions are correct. Otherwise escalate.
8:93 | Detector | Exposure events occurred out-of-order. This typically happens when an error occurs on another node during an exposure. | Check for error conditions on other nodes. Otherwise escalate.
8:94 | Detector | Detector failed to get the temperature set-point information from the DTC. | Check connections to DTC and that DTC is operating correctly. Possible bad DTC hardware.
8:95 | Detector | An unhandled interrupt or processor exception was received. | Check that the correct software and ROS were installed using correct instructions. Otherwise escalate. Logfile Textgram includes details.
8:96 | Detector | An error occurred in the CAN driver. See textgram in logfile for details. | Escalate as possible hardware error.
8:97 | Detector | ROS is incompatible with the Detector hardware and/or software. | Load correct ROS for the Detector.
8:98 | Detector | An unexpected software error occurred. | Reboot Detector and return logfiles to engineering.
8:99 | Detector | Detector has been configured with optional debug and/or diagnostic flags. Issued at start-up. | Clear all debug configuration options from the Detector before returning to customer.
8:100 | Detector | Unsupported modality selected. | Check that correct Detector and readout sequence are loaded. If unresovled, escalate.
8:101 | Detector | Unsupported exposure technique selected. | Check that correct Detector software and readout sequence are loaded. If unresovled, escalate.
8:102 | Detector | Issued when the SerialLite fiber optic link is detected to be down. | Make sure AWS is turned on and the fiber optic cable is securely connected in the correct order. Check PCI firmware and DSS software versions.
8:103 | Detector | There is no valid state machine configuration to support the currently selected modality and technique settings. | Change modality and/or technique. Check that correct Detector software and readout sequence are loaded. If unresovled, escalate.
8:104 | Detector | Array hardware is in unexpected or invalid state. See textgram in logfile for details. | Check that software and firmware versions are properly installed. If unresolved, replace MAC board.
8:105 | Detector | (obsolete)Differential receiver error. | Hardware error. Replace MAC board.
8:106 | Detector | A gate-loop error occurred - either a gate driver ASIC failed, or a connector may not be properly seated. | Check gate FIB connector on MAC and gate FIB power connections. If unresolved, replace Detector.
8:107 | Detector | An internal hardware communication error occurred. See textgram in logfile for details. | Check source FIB seating on MAC board. If unresolved, escalate.
8:108 | Detector | An internal hardware communication error occurred. See textgram in logfile for details. | Check source FIB seating on MAC board. If unresolved, escalate.
8:109 | Detector | An invalid request was received from another node. | Check all nodes for correct software versions and other error conditions. Otherwise escalate.
8:110 | Detector | An error occurred loading the ROS - it is invalid or corrupt. | Check that correct Detector software and ROS are loaded. Otherwise escalate.
8:111 | Detector | An error occurred reading the source FIB temperature sensor. | Possible bad temperature sensor on the source FIB heat sink.
8:112 | Detector | An error occurred configuring the fiber optic frequency synthesizer. | Make sure correct firmware is installed. If unresovled, escalate.
8:113 | Detector | An error occurred in the CPU A/D (conversion error or self-test error). See textgrams in logfile for details. | Escalate as possible hardware error.
8:114 | Detector | An incompatible or unrecognized set of hardware components are present in Detector. | If new components were recently installed, check that they are compatible with this Detector. If unresovled, escalate.
8:115 | Detector | The Detector front chamber heat sink temperature has exceeded the configurable limit. | Check Detector housing for clogged airways and proper fan operation. Check sensor connection to MAC board. If unresovled, escalate.
8:116 | Detector | Detector power failure interrupt was triggered | Check power supply from PMC. If unresolved, escalate.
8:117 | Detector | The installed ROS is using either a method or variable that has been marked as for debug-use only. | Install updated Read-out sequence.
8:118 | Detector | The installed ROS is using either a method or variable that has been marked as obsolete. | None.
8:119 | Detector | The DET issued a request for power data to the PMC (CAN_REQ_DSS_PWR_DATA) and did not hear back (CAN_VAL_DSS_PWR_DATA) within 10 seconds. | Make sure the PMC software is up to date.
8:120 | Detector | An error occurred in the differential receiver (digital phase alignment unlocked). | Hardware error. Replace MAC board.
8:121 | Detector | An error occurred in the differential receiver (channel alignment). | Hardware error. Replace MAC board.
8:122 | Detector | An error occurred in the differential receiver (word alignment). | Hardware error. Replace MAC board.
8:123 | Detector | An error occurred in the differential receiver (CRC). | Hardware error. Replace MAC board.
8:124 | Detector | An error occurred in the differential receiver (framing). | Hardware error. Replace MAC board.
8:125 | Detector | An error occurred in the differential receiver (PLL unlocked). | Hardware error. Replace MAC board.
8:126 | Detector | An error occurred in the differential receiver (FIFO overflow). | Hardware error. Replace MAC board.
8:127 | Detector | ROS timed out waiting for GantryReady event, which exposes a race-condition between ROS and XSM; events-out-of-order error may occur. | Provide logfiles to Engineering.
8:128 | Detector | An error occurred parsing dss.conf or hdt.conf; tomo scan configuration may be incorrect. | Use CalTool to re-configure the projection and binning sections.
12:1 | AWS | Issued when the AWS detects a fault during exposure, or no image data is received from M35. | Examine the logs to determine which fault aborted the exposure.
12:2 | AWS | Issued if the selected view is configured such that it must be exposed without a paddle installed. | Remove the paddle.
12:3 | AWS | The system is configured in High Resolution Tomo mode, and a view containing Tomo was selected, however the AWS is not licensed for High Resolution tomo exposures. | Call service to license the machine for High Resolution.
12:4 | AWS | Missing Detector Configuration -- Issued when the double readout, pixel map, bad line, or weak line files are not found | Calibrate the detector.
12:5 | AWS | Issued when the gain calibration is not found. | Perform gain calibration.
12:6 | AWS | Issued if the installed paddle is not configured as a valid paddle for the currently selected view. | Change the installed paddle to one which the view is configured to support, or select a different view.
12:7 | AWS | The system is configured in 30 projection Tomo mode, and a view containing Tomo was selected, however the AWS is not licensed for 30 projection tomo exposures. | Call service to license the machine for 30 Projections.
12:8 | AWS | Issued when the AWS detects a gantry node reset | Examine the logs to determine which node reset, and why.
12:9 | AWS | Placeholder (not used) | Placeholder (not used)
12:10 | AWS | Placeholder (not used) | Placeholder (not used)
12:11 | AWS | The AWS Image Data Ready monitor did not receive image data or an exposure abort within the configurable amount of time. | Check M35 logging to determine if images were being delivered. Check CAN logs to determine if a node aborted the exposure. Otherwise, check AWS logging to determine the cause of error.
12:12 | AWS | The AWS Exposure Count Monitor did not receive the expected number of image frames within the configurable timeout. | This notification represents the failure of an exposure monitor which has since been deprecated. If this notification is seen, check AWS configuration to ensure that the Exposure Count Monitor is disabled.
12:13 | AWS | The AWS Message Order Monitor received exposure related messages in an unexpected order. | Examine the logs to determine the reason for this error. Messages monitored by this include: CAN_EXP_SCAN_CONFIG_2, CAN_EXP_SCAN_START, CAN_EXP_XRAY_OFF, CAN_EXP_SHOT_DATA, CAN_EXP_SEQUENCE_ABORT and CAN_EXP_SEQUENCE_END.
12:14 | AWS | Issued if the current view does not support mag exposures. | Remove the mag stand, or select a different view.
12:15 | AWS | Invalid or missing AEC calibration. Missing targets, goaltab, dark offsets, or goaltab scale. | Perform AEC calibration on the detector.
12:16 | AWS | Issued when a view requires a geometry calibration and it is not found. Note that this is expected if the system has just been reconfigured from standard resolution to high resolution tomo, or vice versa. | Perform Geometry calibration.
12:17 | AWS | Issued if CAN_EXP_SEQUENCE_END is not received from the GCB | Examine the logs for a reason for this error. If the GCB did send CAN_EXP_SEQUENCE_END, inspect AWS logs.
12:18 | AWS | This verification is performed on startup and is orchestrated by the AWS. It can result in an M35 fault if it fails. | This alert will be cleared when the M35 finishes the wedge verification.
12:19 | AWS | Used to inform the user that the detector is coming up as the AWS is coming up. | Wait for the detector to come online.
12:20 | AWS | Used to inform the user that the detector is unavailable. This may be due to the heartbeat not being sensed. | Wait for the detector to come back online.
12:21 | AWS | A view containing Tomo was selected, however the AWS tomo license is expired. | Call support to license the machine, or select a view which does not contain Tomo.
12:22 | AWS | The compression height is to low for the selected view. In general this will be used for gain calibration to keep the compression device from being imaged | Raise the CDI.
12:23 | AWS | The AWS received an unexpected deny reason of busy. | Examine the log messages to determine what nodes/messages are failing. Verify that AWS Retry On Busy, and other various configuration options are set to appropriate values.
12:24 | AWS | The current view requires the biopsy stage to be installed before an exposure can be taken. | Attach the biopsy stage, or select a non-biopsy view.
12:25 | AWS | The current view cannot be acquired with the biopsy stage. | Detach the biopsy stage, or select a different view.
12:26 | AWS | The appropriate STX or paddle offset calibration cannot be found, or contains invalid data for the current system. | Determine which calibration is missing and perform it. Note that a standard STX calibration will create calibrations for the standard approach. Lateral STX calibration will create STX calibrations for the lateral left and lateral right approaches.
12:27 | AWS | The AWS is not licensed for biopsy related exposures | Call service to license the machine for biopsy, or select a non-biopsy view.
12:28 | AWS | Issued when a biopsy view is selected and no biopsy geometry calibration can be located | Perform the biopsy geometry calibration
12:29 | AWS | Requested during an STX calibration | Move the CARM to the AWS request angle
12:30 | AWS | Requested during an STX calibration | Move the CARM to the AWS request angle
12:31 | AWS | Incorrect needle position for the currently selected view. Needle must be moved prior to exposure. | Move the needle to the correct location
12:32 | AWS | The current paddle position is invalid, based on the AWS config settings for the current view. | Shift the paddle to an allowed position, or change the AWS config for this paddle to allow this position.
12:33 | AWS | The paddle is currently moving horizontally | Wait for the paddle to finish moving before taking an exposure
12:34 | AWS | There is not enough free space on the hard drive to support image capture | Clear hard disk space. Verify that reclaimer configuration settings are set to appropriate values. Verify that the disk is not being filled with SCP images.
12:35 | AWS | The AWS is not licensed for C-View related exposures. | Call service to license the machine for C-View, or select a view which does not contain C-View.
12:36 | AWS | The AWS is not licensed for exposures for this software version. | Call service to license the machine for this software version.
12:37 | AWS | The AWS is not licensed for tomo biopsy related exposures. | Call service to license the machine for tomo biopsy, or select a non-tomo biopsy view.
12:38 | AWS | Tubeoutput.xml is incomplete or contains a mismatched THD serial number. | Redo HVL calibration, or enter correct SN into system tubeoutput.xml file.
12:39 | AWS | The AWS is not licensed for diagnostic related exposures. | Call service to license the machine for diagnostic exposures, or remove diagnostic paddle/mag stand.
12:40 | AWS | Compression height below the minimum height required for tomo reconstructions to occur. | Raise the CDI
12:41 | AWS | The AWS is not licensed for contrast related exposures. | Call service to license the machine for contrast exposures, or select a view which does not contain contrast.
12:42 | AWS | The lateral arm biopsy paddle was installed, but the AWS is not licensed for the lateral arm feature. | Call service to license the system for lateral arm, or remove the lateral arm and lateral paddle.
12:43 | AWS | The AWS failed to parse a minor CAN message. | Call Service.
12:44 | AWS | The AWS failed to parse a major CAN message. | Call Service.
12:45 | AWS | A SmartCurve paddle was installed, but the AWS is not licensed for SmartCurve paddles. | Call service to license the system for SmartCurve paddles, or remove the SmartCurve paddle.
12:46 | AWS | The installed software versions do not match the expected software versions for one or more gantry nodes. | Check the logs to determine which specific nodes are not at the expected version.
14:1 | GCB | test only. Not used | N/A
14:2 | GCB | Issued when an internal software task fails at runtime. | Collect all logs and forward to service. Reload the module software and restore the default settings, recalibration will be required. If the problem persists, replace the board and return the old board for inspection.
14:3 | GCB | Issued when an internal software task fails during init. | Collect all logs and forward to service. Reload the module software and restore the default settings, recalibration will be required. If the problem persists, replace the board and return the old board for inspection.
14:4 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:5 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:6 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:7 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:8 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:9 | GCB | Issued when the VTA has failed to respond in time to a tomo rotation command during the exposure sequence. | Inspect the cable assemblies for proper connections and the VTA mechanical and electrical components. Inspect the log of the exposure sequence for missing, late, or failure responses from the VTA. Check the VTA for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service.
14:10 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:11 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:12 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:13 | GCB | Issued when the ZERO command is required to be issued from the operator to complete configuration. This is expected if the exposure was aborted during a tomo sweep. | If the fault persists after the ZERO is issued, check the tomo pot for correct operation. Recalibration may be required.
14:14 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:15 | GCB | Issued when the BKY cannot retract because of the VTA position at the time of the request. | Inspect the cable assemblies for proper connections and the VTA mechanical and electrical components. Inspect the log of the exposure sequence for missing, late, or failure responses from the VTA. Check the VTA for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:16 | GCB | Issued when the VTA has failed to respond to a rotation command (CARM left/Right or Tube-Arm Left/Right) properly. | Inspect the cable assemblies for proper connections and the VTA mechanical and electrical components. Inspect the log of the exposure sequence for missing, late, or failure responses from the VTA. Check the VTA for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:17 | GCB | Issued when a stuck switch is detected. | Inspect the cable assemblies for proper connections. Check the switch mechanical and electrical components for defects.
14:18 | GCB | Issued when the collimator (THD) is not responding properly to commands or incorrect THD settings detected. | Inspect the cable assemblies for proper connections. Inspect the log of the for missing, late, or failure responses from the THD. Check the THD for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:19 | GCB | Issued when the compression module (CRM) is not responding properly to commands. | Inspect the cable assemblies for proper connections. Inspect the log for missing, late, or failure responses from the CRM and CDI. Check the CRM and CDI mechanical and electrical components for defects or shorts. Check the CRM and CDI for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service.
14:20 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:21 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:22 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:23 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:24 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:25 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:26 | GCB | Issued when an invalid compression paddle ID has been detected. | Remove and reseat the paddle in case the CDI cannot read the RFID tag in its current position. Verify the RFID tag on the paddle has been programmed properly. Verify the CDI tag reader is functioning by installing a known good paddle onto the CDI. If the problem persists, collect all logs and forward to service.
14:27 | GCB | Issued when the CDI has not shifted the paddle to the correct (commanded) position. | Inspect the cable assemblies for proper connections and the paddle position flags. Check the CDI for proper operation, paddle shift calibration may be required. If problem persists, collect all logs and forward to service.
14:28 | GCB | Issued when the BKY has not retracted the grid to the correct (commanded) position. | Inspect the cable assemblies for proper connections. Inspect the log for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:29 | GCB | Issued when the GCB detects an invalid AEC position. | Select a valid AEC position for the current technique. If the problem persists, collect all logs and forward to service.
14:30 | GCB | Issued when invalid technique defaults are detected. They have been automatically re-initialized. | If this persists after reboot, then the GCB may have bad FLASH memory and should be replaced.
14:31 | GCB | Issued when the GCB does not receive the CAN_DET_READY_ON message in time from the DET during an exposure. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence to verify that the DET was armed before the exposure and for missing, late, or failure responses. Check the DET for proper operating mode, modality setting and calibration, recalibration may be required.
14:32 | GCB | Issued when the GCB does not receive the CAN_DET_SCAN_START message in time from the DET during an exposure. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence to verify that the DET was armed before the exposure and for missing, late, or failure responses. Check the DET for proper operating mode, modality setting and calibration, recalibration may be required.
14:33 | GCB | Issued when the GCB does not receive the AEC DATA from the DET in time after the AEC scout exposure. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence to verify that the DET was armed before the exposure and for missing, late, or failure responses. Check the DET for proper operating mode, modality setting and calibration, recalibration may be required.
14:34 | GCB | Issued when the GCB does not receive the CAN_VAL_GEN_ROTOR_READY message in time from the GEN during the exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN. Check the GEN for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:35 | GCB | Issued when the GCB does not receive the CAN_VAL_GEN_XRAY_ON message from the GEN in time. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN or the BKY. Check the GEN and BKY for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:36 | GCB | Issued when the GCB does not receive the post-exposure data from the GEN in time. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN. Check the GEN for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:37 | GCB | Issued when the THD configuration (mirror, filter, aperture) step timed out during an exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the THD. Check the THD for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:38 | GCB | Issued when the CRM configuration (tomo start/zero) step of the exposure sequence times out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the CRM. Check the CRM for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:39 | GCB | Issued when the grid motion configuration step of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:40 | GCB | Issued when the THD filter configuration step of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the THD. Check the THD for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:41 | GCB | Issued when the THD aperture configuration step of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the THD. Check the THD for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:42 | GCB | Issued when the tomo configuration step ( tomo start/zero) of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the CRM. Check the CRM for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:43 | GCB | Issued when the grid movement does not complete correctly during an exposure. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:44 | GCB | Issued when a Tube-Arm rotation error occcurs during an exposure sequence. | Inspect the cable assemblies for proper connections and the VTA rotation and tomo pot circuits. Inspect the log of the exposure sequence for missing, late, or failure responses from the VTA. Check the VTA for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:45 | GCB | Issued when a Tomo motion error occured during an exposure sequence. | Inspect the cable assemblies for proper connections and the VTA rotation and tomo pot circuits. Inspect the log of the exposure sequence for missing, late, or failure responses from the VTA. Check the VTA for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:46 | GCB | Issued when the CAN_VAL_XRAY_ON / OFF sequence from the GEN occurs out of order during an exposure. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN or the BKY. Check the GEN and BKY for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:47 | GCB | Issued when the GCB does not receive the post-exposure data from the GEN during an exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN. Check the GEN for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:48 | GCB | Issued when the AEC shot data is too high and out of range. | Check the current technique settings for anything that may be invalid. Verify the DET AEC parameters, recalibration may be required.
14:49 | GCB | Issued when the AEC calculated mAs is greater than the available mAs for a given technique. | Check the current technique settings for anything that may be invalid. Verify the DET AEC parameters, recalibration may be required.
14:50 | GCB | Issued when the AEC calculated mAs is lower than the available mAs for a given technique. | Check the current technique settings for anything that may be invalid. Verify the DET AEC parameters, recalibration may be required.
14:51 | GCB | Issued when the AEC shot data is too low and out of range. | Check the current technique settings for anything that may be invalid. Verify the DET AEC parameters, recalibration may be required.
14:52 | GCB | Issued when the GCB receives an exposure abort message from another module in the system. | Inspect the log of the exposure sequence for notifications from other modules in the system.
14:53 | GCB | Issued when the GEN boost step of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service. | ROTOR 0x01 THD 0x02 CARM 0x04 TOMO 0x08
14:54 | GCB | Issued when the scan configuration step of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
14:55 | GCB | Issued when a scan does not complete in time during the exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
14:56 | GCB | Issued when the X-ray pulse does not complete properly during an exposure. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
14:57 | GCB | Issued when the post-exposure configuration step of the exposure sequence has timed out. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
14:58 | GCB | Issued when the GEN does not respond to commands from the GCB. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN. Check the GEN for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:59 | GCB | Issued when a scan has timed out during the exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
14:60 | GCB | Issued when the GCB does not receive the AEC or GEN configuration data in time during the exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN or the DET. Check the GEN for proper operation and the system for valid AEC calibration values, recalibration may be required. If the problem persists, collect all logs and forward to service.
14:61 | GCB | Issued when the GCB does not receive the CAN_VAL_GEN_XRAY_OFF message from the GEN in time. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the GEN or the BKY. Check the GEN and BKY for proper operation, recalibration may be required. Check the system for valid AEC calibration values, recalibration may be required. If problem persists, collect all logs and forward to service.
14:62 | GCB | Issued when the X-ray switch is released before the exposure sequence is complete. | Ether operator error (early release) or faulty switch circuitry. Inspect switch hardware for faults or shorts. If problem persists, collect all logs and forward to service.
14:63 | GCB | Issued when there is an invalid C-Arm Zero-detect reading. | Inspect the cable assemblies for proper connections and the zero-sense circuit for defects. Check the CRM for proper operation, recalibration may be required. If problem persists, collect all logs and forward to service.
14:64 | GCB | Issued when the filament calibration shot limit has been reached or if the filament is not calibrated within the allowed shot limit. | This is normal if the filament is way out of calibration, but should only fail once for a particular mA station during calibration. If this error persists, then there may be a legitimate GEN hardware problem and the GEN should be inspected for possible further action.
14:65 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:66 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:67 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:68 | GCB | SPARE - No Longer Used | SPARE - No Longer Used
14:69 | GCB | Issued when the compression height is restricted by the current biopsy needle position. | This is done so the paddle will not hit the biopsy needle. Raise the needle if necessary.
14:70 | GCB | Issued when the BGM is detected but the BCM is not. Self-clearing ALERT. | Connect the BCM. If the BCM is connected, check the cable and connectors for defects. If the BCM screen remains dark, inspect the BCM for damage.
14:71 | GCB | Issued when the BGM heartbeat is detected but not the BGM accessory hooks. | Check the BGM locking and latch mechanism.
14:72 | GCB | Issued when the BGM is detected, but not latched. | Check that the BGM latches are set correctly in place.
14:73 | GCB | Issued when manual rotation of the C-Arm is disabled due to the detected compression force. | Release compression. If no compression is applied and lockout is still in effect, then compression force calibration may be invalid.
14:74 | GCB | Issued when the tomo pot and/or tomo Hall-effect sensor indicate a tomo-non-zero position and is preventing C-Arm manual rotation. | Press Tomo ZERO button and try again. If Tomo-zero is still invalid, then a recalibration of the tomo-zero position may be necessary.
14:75 | GCB | Issued when the detected compression force prevents the requested paddle shift. | The compression force must be zero to allow the paddle shift. If paddle is not compressed, then the compression calibration may be invalid.
14:76 | GCB | Issued when a new exposure is attempted without releasing X-Ray switch after the previous exposure. | Inspect the cable assemblies between the AIO and the GCB for proper connections. Inspect the log for missing, late, or failure responses from the AIO. Check the X-Ray switch mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
14:77 | GCB | Issued when the CDI has not shifted the paddle to the correct (commanded) position. | Check CDI for proper operation. Check CDI shift calibration, position flags.
14:78 | GCB | Issued when the grid retract/lock operation has timed out. | Inspect the cable assemblies for proper connections. Inspect the log for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service.
14:79 | GCB | Issued when the MAG stand is installed with an invalid MAG view selected. | Select a valid MAG view, or remove the MAG stand.
14:80 | GCB | Issued when the C-Arm angle exceeds the limit for a Tomo sweep based on the current VTA height. | Cannot attempt a tomo exposure when the C-Arm is rotated beyond +180 or -140. Change the C-Arm angle or raise the VTA height.
14:81 | GCB | Issued when an invalid accessory ID is detected. | Remove and reattach the accessory, as it may be have been installed in the wrong position or improperly. Inspect the accessory sense circuit for defects. If problem persists, collect all logs and forward to service.
14:82 | GCB | Issued when an invalid paddle and mag view combination is detected. | Make sure the paddle is supported for the current MAG veiw and the paddle ID is correctly programmed.
14:83 | GCB | Issued when an invalid paddle or view is detected | The detected paddle cannot be used with the selected exam. Change to a supported paddle or check that the paddle ID is correctly programmed.
14:84 | GCB | Issued when the face shield is retracted. | Move the face shield to the proper position.
14:85 | GCB | Issued when an invalid face shield position is detected | Face shield must be locked in either in 'IN' or the 'Retracted' position
14:86 | GCB | Issued when the paddle shift (in override mode) does not match view. | Use CDI paddle shift buttons to set the correct shift position.
14:87 | GCB | Issued when FAST mode is selected for a non-FAST paddle (MAG). | Change the paddle or change the FAST setting. Recalibration of the FAST selection switch may be required.
14:88 | GCB | Issued if the paddle height is greater than configured height limit for tomo exposures. | Compression height must be lower than configured height limit. Lower the paddle to continue.
14:89 | GCB | Issued when force is detected but no paddle or valid paddle ID detected. | If there is a paddle installed, check that the paddle RFID tag is programmed correctly and check the CDI paddle reader circuitry. If no paddle is installed, check to see if the CDI is making contact with anything in the system and adjust as necessary. If problem persists, collect all logs and forward to service.
14:90 | GCB | Issued when the GCB detects a fault in the CAN receiver hardware. | Possible bad CAN bus wiring/termination or bad GCB hardware. Inspect the cable assemblies for proper connections and CAN bus termination. If the problem persists, replace the GCB board.
14:91 | GCB | Issued when the GCB has disabled grid movement because both the DET and DTC heartbeats are missing. | Check for proper operation of the DET and DTC. This check can be bypassed by setting the override on the GCB (Utility switch #31). If the problem persists, collect all logs and forward to service.
14:92 | GCB | Issued when the GCB detects an active x-ray switch circuit at X-Ray Button B (RJ14-11) | Check for stuck switch. Inspect the cable assemblies between the AIO and the GCB for proper connections. Check the AIO board mechanical and electrical components for defects or shorts.
14:93 | GCB | Issued when the Biopsy accessory is installed, but a non-biopsy modality is selected. | Select a valid Biopsy modality to continue.
14:94 | GCB | Issued when the tubehead is not at a valid Stereo Position (Center, Left, Right). | Re-zero the Tube-arm. If that fails, then there is a possible hardware positioning problem.
14:95 | GCB | Issued when the Biopsy modality is selected with a CENTER view and manual rotation mode and the arm is not centered. Self-Clears when the correct position is reached. | Move the Tube-Arm to the 0 (CENTER) position.
14:96 | GCB | Issued when the Biopsy modality is selected with a RIGHT view and manual rotation mode and the arm is not in position. Self-clears when the correct position is reached. | Move the Tube-Arm to the +15 (RIGHT) position.
14:97 | GCB | Issued when the Biopsy modality is selected with a LEFT view and manual rotation mode and the arm is not in position. Self-clears when the correct position is reached. | Move the Tube-Arm to the -15 (LEFT) position.
14:98 | GCB | Issued when the Tomo or biopsy sweep cannot be performed due to the C-ARM angle/height violating the pinch-point limit for the requested sweep. | Change the C-ARM angle or VTA height to allow sufficient room to conduct the sweep.
14:99 | GCB | Issued when a BGM latch fault is detected during the exposure sequence. | Remove and reattach the BGM, as it may be have been installed properly. Inspect the BGM latch circuit for defects. If problem persists, collect all logs and forward to service.
14:100 | GCB | Issued when the aperture size is too small for a proper biopsy image. | Increase the aperture size or decrease the compression height.
14:101 | GCB | Issued when the GCB is unable to find valid configuration data stored in FLASH. The data may be corrupt or missing. | Reload the site specific configuration data if available, or restore defaults. Recalibration may be required.
14:102 | GCB | Issued when the heartbeat from the BCM is present but there is no heartbeat from BGM. | Check that the cable from the Affirm to the C-Arm is properly connected. Remove and reseat the cables.
14:103 | GCB | Issued when the generator fan has failed and MUST be serviced before the system can continue. | Service the fan and replace if necessary.
14:104 | GCB | Issued when the GCB does not receive multiple VTA heartbeats in a row over a preset period of time and shutdown the gantry as a precaution. | Collect all logs and forward to service. Inspect the cable assemblies for proper connections and the VTA mechanical and electrical components. Restart the gantry and verify VTA operation and motion.
14:105 | GCB | Issued when the CAN_VAL_GRD_HVEN_ON message was received from the BKY, but no CAN_VAL_XRAY_ON on message was received from the GEN. | Inspect the cable assemblies for proper connections. The cable between the BKY and the GEN used to trigger the GEN may be faulty or connected improperly.
14:106 | GCB | Issued when the CAN_VAL_XRAY_OFF message was received from the GEN, but no CAN_VAL_GRD_HVEN_OFF message was received from the BKY. | Inspect the cable assemblies for proper connections. The cable between the BKY and the GEN used to trigger the GEN may be faulty or connected improperly.
14:107 | GCB | Issued when the VTA has failed to respond in time to a query if tomo rotation is allowed. | Inspect the cable assemblies for proper connections and the VTA mechanical and electrical components. Inspect the log of the exposure sequence for missing, late, or failure responses from the VTA. Check the VTA for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service.
16:1 | GSW_01 | Issued when a stuck switch has been detected | Check for stuck switch. Inspect all housing components proper fit and connection. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
17:1 | GSW_02 | Issued when a stuck switch has been detected | Check for stuck switch. Inspect all housing components proper fit and connection. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
25:1 | HV Generator | Issued when the 600V rail voltage is out of tolerance. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:2 | HV Generator | Issued when the ADC busy signal does not go inactive. | If problem persists collect all logs and forward to service, board replacement may be necessary.
25:3 | HV Generator | Issued when the Generator Control board temperature has exceeded the limits | Check line voltages and on-board regulator. Replace board if problem persists.
25:4 | HV Generator | Issued when the digital POT setup fails on the GEN control board. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:5 | HV Generator | Issued when the exhaust fan has been issuing an alert for over 10 seconds. | Check line voltages and on-board regulator. Inspect the fan and all connections. Replace the fan if necessary.
25:6 | HV Generator | Issued when the exposure time parameter is out of range. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
25:7 | HV Generator | Issued when the GEN receives an unexpected or unhandled error from the onboard CPLD. | Check for GEN and CPLD software compatibility, reload as necessary. If problem persists collect all logs and forward to service, board replacement may be necessary.
25:8 | HV Generator | Issued when the filament current varies by more than 12.5% of the target during boost. | Filament calibration is required. If the problem persists, check filament circuit and HV chain for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:9 | HV Generator | Issued when the CPLD backup timer circuit is bypassed. (CPLD DIP switch 4 is ON). This is only allowed in NPT mode. Function is only used for HIPOT. | Disable the backup timer circuit bypass or activate NPT mode.
25:10 | HV Generator | Issued when the GEN control board detects a Filament over current condition | Possible calibration issue, redo all calibrations first. Check filament and HV circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:11 | HV Generator | Issued when the GEN control board detects a filament over voltage condition. | Filament calibration is required. If the problem persists, check filament circuit and HV chain for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:12 | HV Generator | Issued when a filament starting current is found out of range in EEPROM. | Possible calibration issue, redo all calibrations first. Check filament and HV circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:13 | HV Generator | Issued when the grid bias sense exceeds the 5% tolerance. | Possible calibration issue, redo all calibrations first. Check filament and HV circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:14 | HV Generator | Issued when the heat sink fan has been issuing an alert for over 10 seconds. | Check line voltages and on-board regulator. Inspect the fan and all connections. Replace the fan if necessary. Check all components in HV chain for defects or shorts.
25:15 | HV Generator | Issued when the GEN control board detects a High Voltage Arc condition | Intermittent Arcing: can happen when tube is very hot. Persistent Arcing: check HV connections and HV grease. Check all components in HV chain.
25:16 | HV Generator | Issued when the GEN heat sink temperature has exceeded the allowable limit. | Check all HV connections and HV grease. Check all components in HV chain for defects or shorts.
25:17 | HV Generator | Issued when the GEN control board detects an inverter fault condition. | Check all HV connections and HV grease. Check all components in HV chain for defects or shorts.
25:18 | HV Generator | Issued when the GEN control board detects an inverter over current condition. | Check all HV connections and HV grease. Check all components in HV chain for defects or shorts.
25:19 | HV Generator | Issued when the -15V rail voltage is out of tolerance. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:20 | HV Generator | Issued when an invalid exposure parameter is received. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
25:21 | HV Generator | Issued when the +15V rail voltage is out of tolerance. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:22 | HV Generator | Issued when the +5V bus voltage on the GEN controller board is out of tolerance. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:23 | HV Generator | Issued when the +5V Ref voltage on the GEN controller board is out of tolerance. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:24 | HV Generator | (obs)Issued when the Rotor Brake function produces an error | Check line voltages and on-board regulator. Check rotor brake circuit. Possibly bad tube.
25:25 | HV Generator | Issued when the GEN control board detects a rotor fault condition. | Check line voltages and on-board regulator. Check rotor and brake circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:26 | HV Generator | Issued when the rotor current sense is found below threshold during boost. | Check line voltages and on-board regulator. Check rotor and brake circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:27 | HV Generator | Issued when the rotor ready detection times out after boost is initiated. | Check line voltages and on-board regulator. Check rotor and brake circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:28 | HV Generator | (obs)Issued when the GEN control board detects a Rotor Over Current condition | Check line voltages and on-board regulator. Check rotor and brake circuit. Possibly bad tube.
25:29 | HV Generator | (obs)Issued when the GEN control board detects a Rotor System fault condition | Check line voltages and on-board regulator. Check rotor and brake circuit. Possibly bad tube.
25:30 | HV Generator | Issued when the GEN control board detects an X-Ray tube over current condition. | Check the loading and temperature for signs of an overheated X-Ray tube, rest tube as necessary. Possible calibration issue, redo all filament calibrations. Check the filament drive circuit for defects or shorts. If the problem persists board replacement may be necessary.
25:31 | HV Generator | Issued when the GEN control board detects an X-Ray tube over voltage condition. | Check the loading and temperature for signs of an overheated X-Ray tube, rest tube as necessary. Possible calibration issue, redo all filament calibrations. Check the filament drive circuit for defects or shorts. If the problem persists board replacement may be necessary.
25:32 | HV Generator | (obs)Issued when an unexpected and unhandled message is received at the msg router | Incompatible software versions
25:33 | HV Generator | Issued when the Plus 15V rail is measured and found out of spec | Check line voltages and on-board regulator. Replace board if problem persists.
25:34 | HV Generator | Issued when the maximum exposure count has been exceeded during this exposure sequence. | If this error occurred during filament calibration clear the fault and take a second exposure to complete the current station. Check current modality settings if observed outside the calibration task. If the problem persists, collect all logs and forward to service.
25:35 | HV Generator | Issued when an invalid result is calculated from GEN calibration data. | Possible bad calibration data, reload defaults and recalibrate. If the problem persists then possible bad EEPROM of filament circuit, inspect board for defects or shorts. Board replacement may be necessary.
25:36 | HV Generator | Issued when an auto filament calibration routine failed to adjust the filament enough to pass calibration. | Possible bad calibration data, reload defaults and recalibrate. If the problem persists then possible bad EEPROM of filament circuit, inspect board for defects or shorts. Board replacement may be necessary.
25:37 | HV Generator | Issued when an exposure is aborted due to the rotor failing to go ready. | Check line voltages and on-board regulator. Check rotor and brake circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:38 | HV Generator | Issued when an exposure is aborted due to a failure in the exposure interlock circuit. | Check the interlock circuit for defects or shorts. Check all wiring and connections. If the problem persists board may be necessary.
25:39 | HV Generator | Issued when a rail fuse sense fault is detected. | Inspect the fuse and sense circuits for defects or shorts. Replace as necessary.
25:40 | HV Generator | Issued when any of the high-voltage feedback connectors are disconnected | Check line voltages, HV feedback connectors, and on-board regulator. Check rotor and brake circuit. Possibly bad tube.
25:41 | HV Generator | Issued when the +15V UV voltage is out of tolerance. | Check line voltages and on-board regulators. If the problem persists board replacement may be necessary.
25:42 | HV Generator | Issued when the exhaust fan fuse sense is in a fault state. | Inspect the fuse and sense circuit for defects or shorts. Replace if necessary.
25:43 | HV Generator | Issued when the heat sink fan fuse sense is in a fault state. | Inspect the fuse and sense circuit for defects or shorts. Replace if necessary.
25:44 | HV Generator | Issued when the rotor is not started due to other fault conditions. | Clear or resolve any existing fault conditions. If the problem persists, collect all logs and forward to service.
25:45 | HV Generator | (obs)Issued when next exposure is prevented, usually when Xray switch not released. | should only happen during bench test
25:46 | HV Generator | Issued when calibration or the adaptive filament adjustments have reached the maximum allowable limits. | Possible calibration issue, redo all calibrations first. Inspect the X-Ray tube for signs of degraded performance. If the problem persists tube replacement may be necessary.
25:47 | HV Generator | Issued when the hardware backup timer times out. | Check for compatible CPLD version and for correct CPLD DIP switch settings. Collect all logs and forward to service as soon as possible.
25:48 | HV Generator | Issued when the X-ray switch is released before the exposure sequence is complete. | Verify any operator errors, inspecting the logs if necessary. Check for defective X-Ray switch circuits.
25:49 | HV Generator | Issued when the HV Shut Down terminated the exposure. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:50 | HV Generator | Issued if the exposure was aborted before the ramp due to a control interlock error. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:51 | HV Generator | Issued if the exposure was aborted after the ramp due to a control interlock error. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:52 | HV Generator | Issued when the exposure is aborted because no starting mA had been acquired. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:53 | HV Generator | Issued when an exposure is aborted due to an uncleared fault. | Should only happen during bench-test. Otherwise collect all logs and forward to service.
25:54 | HV Generator | Issued when the exposure is aborted inside the control loop. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:55 | HV Generator | Issued when the kV or mA sense exceeds the off limit prior to boost. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:56 | HV Generator | Issued when the starting mA exceeds tolerance at the start of the exposure. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:57 | HV Generator | Issued when the kV sense exceeds tolerance at the start of the exposure. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:58 | HV Generator | Issued when the rotor current sense exceeded the off limit value. | Possible calibration issue, redo all calibrations first. Check filament circuit for defects or shorts. Inspect the X-Ray tube for signs of degraded performance. If the problem persists board or tube replacement may be necessary.
25:59 | HV Generator | Issued when the HVEN signal is asserted when not in an active exposure sequence. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service.
25:60 | HV Generator | Issued when the adaptive filament values have drifted too far from the original calibrated values. | Perform a filament calibration. | mA station | dacDiff | int16_t | DAC_Diff
25:61 | HV Generator | Issued when the X-Ray tube is showing signs of end-of-life (tube output degradation) | Inspect the X-Ray tube for degraded performance and replace if necessary. | current DAC reading
25:62 | HV Generator | Issued when a rotor fuse fault has been detected. | Inspect the rotor fuse and replace if necessary.
25:63 | HV Generator | Issued when a fault has been detected with the fast bleeder fuse. | Inspect the fuse circuit and replace the fast bleeder fuse if necessary.
25:64 | HV Generator | Issued when a fault has been detected by the 1.8V monitor circuit. | Check all voltages and on-board regulators. Check the electrical components for defects or shorts. Board replacement may be necessary.
25:65 | HV Generator | Issued when a grid (BKY) exposure was initiated, but the HVEN signal was not received by the GEN. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service.
25:66 | HV Generator | Issued when a grid (BKY) exposure was initiated, but the HVEN signal to the GEN terminated prematurely. | Inspect the cable assemblies for proper connections. Inspect the log of the exposure sequence for missing, late, or failure responses from the BKY. Check the BKY for proper operation, recalibration may be required. If the problem persists, collect all logs and forward to service. | elapsed time
26:1 | Biopsy Control | Issued when the calculated checksum of the image files do not match the stored value. | Check for compatible software versions and reinstall the matching BCM image files. If the problem persists, collect all logs and forward to service.
26:2 | Biopsy Control | Image files stored in flash memory are incompatible with currently installed software version. | Install correct version of screen images and/or software.
26:3 | Biopsy Control | Issued whan an internal software error has been detected. See error code. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
26:4 | Biopsy Control | Issued when a software task has reported a stack overflow. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
26:5 | Biopsy Control | Issued when an unexpected interrupt or processor exception has occurred. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
26:6 | Biopsy Control | Issued when an error has been detected in the CAN driver software. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
26:7 | Biopsy Control | Issued when the BCM failed POST. | Check board and cabling. Review logs for specific POST failure and address as necessary. If the problem persists, collect all logs and forward to service.
26:8 | Biopsy Control | Issued when the JOG target is out of permissible JOG range. | Move target or switch to Target Guidance Screen. Check calibrations, as recalibration may be required. If the problem persists, collect all logs and forward to service.
26:9 | Biopsy Control | Issued when the BGM does not respond to a request for the current position projected to the target Z plane. | Check boards, cabling, and for compatible software versions. If the problem persists, collect all logs and forward to service.
26:10 | Biopsy Control | Issued when the calculated checksum of the BCM application code does not match the stored value. | Check for compatible software versions and reinstall the BCM software. If the problem persists, collect all logs and forward to service.
26:11 | Biopsy Control | Issued when the calculated checksum of the BCM bootloader code does not match the stored value. | Check for compatible software versions and reinstall the BCM bootlloader. If the problem persists, collect all logs and forward to service.
26:12 | Biopsy Control | Issued when an invalid target, or a target not in the current target list, has been selected. | Select a valid target. If the problem persists, collect all logs and forward to service.
26:13 | Biopsy Control | Issued when motion times out while waiting for the BGM to issue the MOTION_COMPLETE message. | Check boards, cabling, and calibrations. Recalibration may be required. If the problem persists, collect all logs and forward to service.
26:14 | Biopsy Control | Issued when the calculated checksum of the string table does not match the stored value | Check for compatible software versions and reinstall the matching BCM string tables. If the problem persists, collect all logs and forward to service.
26:15 | Biopsy Control | Issued when an error occurrs while attempting to erase or write to an area flash memory. | Re-attempt the software installation or configuration update. If the problem persists, collect all logs and forward to service.
26:16 | Biopsy Control | Issued when the string table stored in flash memory is incompatible with the currently installed software. | Check for compatible software versions and reinstall the matching BCM string tables. If the problem persists, collect all logs and forward to service.
26:17 | Biopsy Control | Issued when in lateral mode and the C-Arm is not at zero degrees. | Move the C-Arm to zero degrees. If at zero degrees and the problem persists, collect all logs and forward to service.
26:18 | Biopsy Control | Issued when in lateral mode and a left or right approach has not been specified. | Set the left or right approach on the BCM.
26:19 | Biopsy Control | Issued when in lateral mode and the lateral offset stand has not been verified to be in place. | Select OK to verify stand is in place on the BCM.
26:20 | Biopsy Control | Issued when an invalid lateral mode setting has been detected. | Check for compatible software versions. If the problem persists, collect all logs and forward to service. | bad data
26:21 | Biopsy Control | Issued when the lateral arm X value for a target is outside the valid range. | Check target, move as necessary, and try again. Recalibration may be required. If the problem persists, collect all logs and forward to service.
27:1 | Tubehead | Issued when the calculated new motor position is out of range. | Recalibrate the board. If the problem persists, collect all logs and forward to service.
27:2 | Tubehead | Issued when an invalid aperture calibration position parameter has been received. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:3 | Tubehead | Issued when an invalid aperture motion position parameter has been received. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:4 | Tubehead | Issued when a blade incremental move command fails. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:5 | Tubehead | Issued when a blade absolute move command fails. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:6 | Tubehead | (obs)issued when the THD receives an invalid position parameter for filter calibration | Cal Tool bug, or incompatible software versions.
27:7 | Tubehead | Issued when an invalid filter motion position parameter has been received. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:8 | Tubehead | Issued when an invalid lamp calibration position parameter has been received. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:9 | Tubehead | Issued when an invalid lamp motion position parameter has been received. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
27:10 | Tubehead | Issued when POST fails to position a motor within the accuracy bounds. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:11 | Tubehead | Issued when POST fails to home all motors. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:12 | Tubehead | Issued when POST fails to position a motor and times out. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:13 | Tubehead | (obs)Test error for THD- new device
27:14 | Tubehead | (obs)Issued when an unexpected and unhandled message is received at the msg router | incompatible software versions.
27:15 | Tubehead | issued when the THD fails to execute a motion command for any reason | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:16 | Tubehead | Issued when the filter home command fails. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:17 | Tubehead | Issued when a board voltage is out of tolerance. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
27:18 | Tubehead | Issued when the tube fan has failed. | Check fan, line voltages and on-board regulators. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
27:19 | Tubehead | Issued when the lamp fan has failed. | Check lamp, fan, line voltages and on-board regulators. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
27:20 | Tubehead | Issued when the filter does not reach the commanded position. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:21 | Tubehead | Issued when one or more of the motors commanded to go home fail to go home. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:22 | Tubehead | Issued when a motor accuracy failure has occurred. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:23 | Tubehead | Issued when the lamp mirror assembly fails to position properly. | Possible mechanical problems. Check calibrations and inspect all mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:24 | Tubehead | Issued when a motor encoder counter was unexpectedly cleared. Possibly the result of a transient on the home flag. | When this fault is detected, all of the DC motors are homed and the fault can be cleared. If the problem persists, collect all logs and forward to service.
27:25 | Tubehead | Issued when multiple filter-wheel positioning problems have been detected. | Check calibrations and alignments. Inspect all sensors, encoders and mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:26 | Tubehead | Issued when multiple aperture blade positioning problems have been detected. | Check calibrations and alignments. Inspect all sensors, encoders and mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:27 | Tubehead | Issued when un-commanded filter wheel motion has been detected. | Check calibrations and alignments. Inspect all sensors, encoders and mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
27:28 | Tubehead | Issued when filter recovery has failed. Possibly after un-commanded motion. | Check calibrations and alignments. Inspect all sensors, encoders and mechanical components for defects. Recalibration or part replacement may be required. If the problem persists, collect all logs and forward to service.
29:1 | VTA | Issued when a message received contained data that was either too large or too small to be processed. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
29:2 | VTA | Issued when the directed motion is either too large or too small to complete. | Check the source of the motion command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
29:3 | VTA | (obs)The error message will contain the unrecognized ID | This a software bug in the VTA or Cal Tool (whichever node sent the CMD), or possibly incompatible software versions.
29:4 | VTA | issued when POST detects a problem with the CARM rotation transducer. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:5 | VTA | Issued when POST detects a problem with the TOMO rotation transducer. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:6 | VTA | Issued when POST detects a problem with the Vertical travel transducer. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:7 | VTA | issued when POST detects a problem with the CARM rotation calibration. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:8 | VTA | Issued when POST detects a problem with the TOMO rotation calibration. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:9 | VTA | Issued when POST detects a problem with the Vertical travel calibration. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:10 | VTA | Issued when manual rotation is attempted with an unresolved POST fault. | Investigate the original POST fault.
29:11 | VTA | Issued when TOMO motion is attempted with an unresolved POST fault | Investigate the original POST fault.
29:12 | VTA | Issued when TOMO motion is attempted with the grid not retracted | Investigate the original POST fault.
29:13 | VTA | Issued when vertical motion is attempted with an unresolved POST fault | Investigate the original POST fault.
29:14 | VTA | Issued when the measured CARM/Tomo motion is slow or in the wrong direction | Inspect the drive belt and linkage for slippage or other defects. Perform all mechanical alignment and motion calibrations. Inspect the drive circuit for any defects or shorts. If the problem persists, collect all logs and forward to service.
29:15 | VTA | Issued when uncommanded C-Arm Motion was detected. | Check the C-Arm pot connections and for any defects or shorts. Check cables for insulation wear and shorts. Inspect the drive circuit for any defects or shorts. Board or part replacement may be required. If the problem persists, collect all logs and forward to service. | final C-Arm angle
29:16 | VTA | Issued when uncommanded Tomo Motion was detected. | Check the tomo pot connections and for any defects or shorts. Check and adjust the rotation worm gear backlash as necessary. Inspect the drive circuit for any defects or shorts. Board or part replacement may be required. If the problem persists, collect all logs and forward to service. | final tomo angle
29:17 | VTA | Issued when uncommanded vertical motion has been detected or when the vertical motion has stalled. | Check the vertical travel pot connections and for any defects or shorts. Check the vertical travel gear belt and adjust as necessary as a loose belt can cause the pot to slip. Inspect the drive circuit for any defects or shorts. Board or part replacement may be required. If the problem persists, collect all logs and forward to service.
29:18 | VTA | Issued when the 24V voltage check has failed. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:19 | VTA | Issued when the 15V voltage check has failed. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:20 | VTA | Issued when the 12V voltage check has failed. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:21 | VTA | Issued when the 5V voltage check has failed. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:22 | VTA | Issued when the VCC voltage check has failed. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:23 | VTA | Issued when the voltage on the 24V line has moved outside the Tier 2 voltage. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:24 | VTA | Issued when the 24V line has moved outside the Tier 1 voltage. | Check all line voltages feeding the VTA. Inspect connections and on-board regulator for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
29:25 | VTA | Issued when a fault has occurred when reading the Internal A/D ports. | If the problem persists, collect all logs and forward to service as board replacement may be required.
29:26 | VTA | Issued when the VTA exhibits uneven movement. | Possible mechanical problems. Inspect all mechanical components for defects and perform all mechanical alignment and motion calibrations. If the problem persists, collect all logs and forward to service.
29:27 | VTA | Issued when the vertical POT resolution has more than a 6% deviation from the nominal value. | Check the VTA height limits calibration and recalibrate if necessary. Check the vertical travel pot connections and for any defects or shorts. If the problem persists, collect all logs and forward to service.
29:28 | VTA | Issued when the system is configured for only 2D operation and Tomo motion is not allowed | Check system for license and for compatible software in all modules.
29:29 | VTA | Issued when noise is detected on the tomo potentiometer channel. | Check the tomo potentiometer connections and the potentiometer itself for defects or shorts. If the problem persists, replace the tomo potentiometer. | Expected speed (delta profile)
29:30 | VTA | Issued when noise is detected on the C-Arm potentiometer ADC channel. | Check the C-Arm potentiometer connection and the C-Arm potentiometer for defects or shorts. If problem persists, collect all logs and forward to service. C-Arm potentiometer replacement may be necessary. | Expected speed (delta profile)
29:31 | VTA | Issued when noise is detected on the C-Arm/Tomo potentiometer ADC channel. | Check the C-Arm/Tomo potentiometer connections and the potentiometer itself for defects or shorts. If the problem persists, replace the C-Arm/Tomo potentiometer.
29:32 | VTA | Issued when the C-Arm is too low and unable to reach requested angle. | Please raise the C-Arm. Recalibration may be required. If problem persists, collect all logs and forward to service.
31:1 | Bucky | (obs)Issued when grid motion is commanded before being set up | If during exposure, this is a GCB software bug. Else CalTool bug.
31:2 | Bucky | Issued when grid motion has timed out before completion. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. If the problem persists, collect all logs and forward to service.
31:3 | Bucky | (obs)Issued when retract motor motion timed out before completion | Check mechanics and drive electronics.
31:4 | Bucky | Issued when motion has been aborted because the home flag was not detected. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. If the problem persists, collect all logs and forward to service.
31:5 | Bucky | Issued when a command is received with an invalid parameter. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
31:6 | Bucky | Issued when a grid retract command is received when the grid was in motion or not in the home position. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. If the problem persists, collect all logs and forward to service.
31:7 | Bucky | Issued when a command to move the grid is received with the C-Arm not zeroed. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
31:8 | Bucky | Issued when the 3.3VDC bus voltage is out of tolerance. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
31:9 | Bucky | Issued when the 5VDC bus voltage is out of tolerance. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
31:10 | Bucky | Issued when the VBB1 bus voltage is out of tolerance. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
31:11 | Bucky | The VBB2 VDC bus on the BKY board is out of tolerance - (obs)issued when a DC bus on the board is out of tolerance | Check line voltages and on-board regulator.
31:12 | Bucky | Issued when both home flags are active at the same time. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. If the problem persists, collect all logs and forward to service.
31:13 | Bucky | Issued when a grid home command fails. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. If the problem persists, collect all logs and forward to service.
31:14 | Bucky | Issued when the HVEN ON signal from the CPLD has timed out. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. If the problem persists, collect all logs and forward to service.
31:15 | Bucky | Issued when a motion fault has occurred and the grid is executing a 16-sec recovery sequence. This should self-clear in approximately 16 seconds. | Possible mechanical or electrical problems if fault does not self clear. Inspect all mechanical and electrical components for defects and recalibrate. The grid drive mechanism or grid device may require replacing. If the problem persists, collect all logs and forward to service.
31:16 | Bucky | Issued when the CPLD does not generate the X-Ray On/Off signal during a 2D exposure. GCB/DET notifications may also be issued in relation to this failure. | Check the BKY CPLD version. The BKY CPLD may need to be upgraded. If the problem persists, collect all logs and forward to service. | XRAY_ON or XRAY_OFF encoder counts, depending on the information in byte[4]
31:17 | Bucky | Issued when the grid drifts away from home, particularly during the exposure routine. | This issue may coincide with reports of gridlines in images. Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. The grid drive mechanism or grid device may require replacing. If the problem persists, collect all logs and forward to service.
31:18 | Bucky | Issued when the grid drifts away from home and the homing routine failed to return it. | Possible mechanical or electrical problems. Inspect all mechanical and electrical components for defects and recalibrate. The grid drive mechanism or grid device may require replacing. If the problem persists, collect all logs and forward to service.
31:19 | Bucky | Issued when the grid does not retract before the timeout has expired. | Possible mechanical problems. Inspect all mechanical components for defects and recalibrate. The grid motor may require replacing. If the problem persists, collect all logs and forward to service. | time it took to retract in msx10 | retract_mean | uint16_t | retract time running average for this type
31:20 | Bucky | Issued when the encoder counts did not change during grid scan. GCB/DET notifications may also be issued in relation to this failure. | Check the BKY CPLD version. The BKY CPLD may need to be upgraded. Check the home sensors. If the problem persists, collect all logs and forward to service.
32:1 | C-arm | (obs)Issued when an unexpected and unhandled message is received at the msg router | Incompatible software versions
32:2 | C-arm | (obs)Issued when VTA detects a software fault (during development) | Incompatible software versions
32:3 | C-arm | (OBS)The 3.3 VDC bus on the CRM board is out of tolerance issued when a DC bus on the board is out of tolerance | check line voltages and on-board regulator.
32:4 | C-arm | Issued when the 12VDC bus voltage is out of tolerance. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
32:5 | C-arm | Issued when the 5VDC bus voltage is out of tolerance. | Check line voltages and on-board regulator. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
32:6 | C-arm | Issued when a manual motion command timed out or stopped without a stop command. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
32:7 | C-arm | Issued when a bad parameter has been received by the C-Arm in a set compression force command. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
32:8 | C-arm | Issued when a bad parameter has been received in a set compression thickness command. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
32:9 | C-arm | Issued when the Zero detect values indicate a short or open circuit. | Inspect the zero-detect electronics (Hall effect sensors) for defects or shorts. If problem persists, collect all logs and forward to service.
32:10 | C-arm | Issued when a fault has been detected with one or more of the array fans. | Check line voltages and on-board regulator. Inspect the fan and all connections. Replace the fan if necessary.
32:11 | C-arm | Issued when the CRM detects a stuck switch during boot. | Inspect the switch module and I2C bus and connections for defects or shorts. If problem persists, collect all logs and forward to service.
32:12 | C-arm | Issued when the CRM has detected an invalid code from the switch module. | Inspect the switch module and I2C bus and connections for defects or shorts. If problem persists, collect all logs and forward to service.
32:13 | C-arm | Issued when sudden loss of compression force is detected. | Check for broken paddle. Inspect brake and clutch assembly for mechanical defects. Replace parts if necessary.
32:14 | C-arm | Issued when attempting to set pre or full compression parameters out of the valid range. | Check the source of the command, including CalTool if applicable, and verify all versions of software are compatible. If the problem persists, collect all logs and forward to service.
32:15 | C-arm | Issued when a fault has been detected in the accessory detection circuit I2C controller. | Self-clears if re-sync is successful or on a gantry power cycle. If the problem persists, collect all logs and forward to service.
32:16 | C-arm | Issued when the left C-Arm switch has reported as stuck. | Check for stuck switch or I2C bus lockup. Inspect all switch bank components proper fit and connections. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
32:17 | C-arm | Issued when the right C-Arm switch has reported as stuck. | Check for stuck switch or I2C bus lockup. Inspect all switch bank components proper fit and connections. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
32:18 | C-arm | Issued when the center C-Arm switch has reported as stuck. | Check for stuck switch or I2C bus lockup. Inspect all switch bank components proper fit and connections. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
32:19 | C-arm | Issued when the left C-Arm switch has reported as stuck at runtime. | Check for stuck switch or I2C bus lockup. Inspect all switch bank components proper fit and connections. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
32:20 | C-arm | Issued when the right C-Arm switch reports as stuck at runtime. | Check for stuck switch or I2C bus lockup. Inspect all switch bank components proper fit and connections. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
32:21 | C-arm | Issued when the center C-Arm switch has reported as stuck at runtime. | Check for stuck switch or I2C bus lockup. Inspect all switch bank components proper fit and connections. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
32:22 | C-arm | Issued when the accessory detection circuit provides invalid readings. | Inspect accessory detection circuit for defects, shorts, or an I2C bus lockup. If problem persists, collect all logs and forward to service.
32:23 | C-arm | Issued when the face shield detection circuit provides invalid readings. | Inspect face shield detection circuit for defects, shorts, or a I2C bus lockup. If problem persists, collect all logs and forward to service.
33:1 | DTC | Issued when the ambient temperature is out of limits | Check room temperature and correct.
33:2 | DTC | Issued when the ambient humidity is out of limits | Check room humidity and correct.
33:3 | DTC | Issued when cooling power has been reduced to avoid internal condensation | Check room temperature and humidity and correct.
33:4 | DTC | Issued for an invalid configuration | Cycle power to load default config. Possible bad DTC PCB if persistent.
33:5 | DTC | Issued when the TEC heat sink temperature is above limit | TEC heat sink is overheating. Possible problem with heat sink fans.
33:6 | DTC | Issued when the coldplate temperature is above limit | Coldplate is overheating. Possible assembly error - check fasteners on coldplate.
33:7 | DTC | Issued when there is a glass temperature sensor fault | Contact TS, possible bad sensor or cable.
33:8 | DTC | Issued when there is a heat sink temperature sensor fault | Contact TS, possible bad sensor or cable.
33:9 | DTC | Issued when there is a coldplate temperature sensor fault | Contact TS, possible bad sensor or cable.
33:10 | DTC | Issued when there is an ambient temperature/humidity sensor fault | Contact TS, possible bad sensor or cable.
33:11 | DTC | Issued when there is a front temperature/humidity sensor fault | Contact TS, possible bad sensor or cable.
33:12 | DTC | Issued when there is an A-to-D converter error | Contact TS, possible bad DTC PCB.
33:13 | DTC | Issued when the 24 volt supply is out of limits | Contact TS, possible bad hardware.
33:14 | DTC | Issued when there is a TEC drive voltage fault | Contact TS, possible bad TEC assembly or bad relay on DTC PCB.
33:15 | DTC | Issued when a hardware fault is sensed | Contact TS, possible bad hardware.
34:1 | Biopsy Guidance | Issued when motor VBB is active while the sense switch is inactive. Condition could allow for uncommanded motion. | Check all motor voltages and cables for defects or shorts. If the problem persists, collect all logs and forward to service as board replacement may be necessary.
34:2 | Biopsy Guidance | Issued when the motor sense switch does not activate VBB. Motor motion wil be prevented. | Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
34:3 | Biopsy Guidance | Issued when board voltages have been detected to be out of tolerance during POST. Condition could cause unpredictable behavior. | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:4 | Biopsy Guidance | Issued when the X-axis motor has failed to position correctly. | Check motor assembly, cables, and POT for defects or shorts. Recalibration may be required. If the problem persists, collect all logs and forward to service as replacement may be necessary.
34:5 | Biopsy Guidance | Issued when the Y-axis motor has failed to position correctly. | Check motor assembly, cables, and POT for defects or shorts. Recalibration may be required. If the problem persists, collect all logs and forward to service as replacement may be necessary.
34:6 | Biopsy Guidance | Issued when the Z-axis motor has failed to position correctly. | Check motor assembly, cables, and POT for defects or shorts. Recalibration may be required. If the problem persists, collect all logs and forward to service as replacement may be necessary.
34:7 | Biopsy Guidance | Issued when a move command was received with an invalid X-coord. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:8 | Biopsy Guidance | Issued when a jog command was received with an invalid X-coord. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:9 | Biopsy Guidance | Issued when a move command was received with an invalid Y-coord. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:10 | Biopsy Guidance | Issued when a move command was received with an invalid Z-coord. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:11 | Biopsy Guidance | Issued when a jog command was received with an invalid Y-coord. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:12 | Biopsy Guidance | Issued when a jog command was received with an invalid Z-coord. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:13 | Biopsy Guidance | Issued when the hardware has reported a motor drive fault. | Inspect the log for other motor faults. Check the motor assembly and cables for defects or shorts. Recalibration may be required. If the problem persists, collect all logs and forward to service as replacement may be necessary.
34:14 | Biopsy Guidance | Issued when the lateral arm has been detected to be in the wrong state. | Check all lateral arm settings. If the problem persists, collect all logs and forward to service.
34:15 | Biopsy Guidance | Issued when the left Affirm latch is not detected to be in the latched state. | Close the latch completely, reseating the Affirm if necessary. Check latch assembly and cables for any mechanical or electrical defects or shorts. Check the board for proper operation. Collect all logs and forward to service. Board replacement may be necessary.
34:16 | Biopsy Guidance | Issued when the right Affirm latch is not detected to be in the latched state. | Close the latch completely, reseating the Affirm if necessary. Check latch assembly and cables for any mechanical or electrical defects or shorts. Check the board for proper operation. Collect all logs and forward to service. Board replacement may be necessary.
34:17 | Biopsy Guidance | Issued when a stuck switch has been detected. | Check for stuck switch. Inspect all housing components proper fit and connection. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service. | state
34:18 | Biopsy Guidance | Issued when the STX matrix transfer from the AWS has failed due to a transfer time out. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:19 | Biopsy Guidance | Issued when a matrix transfered from the AWS has failed the checksum validation. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:20 | Biopsy Guidance | Issued when a guidance hardware calibration is required. | Perform all necessary hardware guidance calibrations. If the problem persists, collect all logs and forward to service.
34:21 | Biopsy Guidance | Issued when a guidance space calibration (STX) is required. | Perform all necessary STX calibrations. If the problem persists, collect all logs and forward to service.
34:22 | Biopsy Guidance | Issued when a safety margin calibration is required. | Perform all required safety margin calibrations. If the problem persists, collect all logs and forward to service.
34:23 | Biopsy Guidance | Issued when the BGM requires an STX matrix. | Transfer a valid STX matrix structure to the BGM. STX calibration may be required. If problem persists, collect all logs and forward to service.
34:24 | Biopsy Guidance | Issued when an attempt is made to add a new DET serial number to a full list. | Clear at least one DET serial number in the list in order to proceed.
34:25 | Biopsy Guidance | Issued when the 10bit ADC failed to complete a conversion. | Check all line voltages and grounds on the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:26 | Biopsy Guidance | Issued when a 10bit ADC conversion FIFO has overrun itself. | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:27 | Biopsy Guidance | Issued when the 15V bus voltage is out of tolerance. (High) | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:28 | Biopsy Guidance | Issued when the 15V bus voltage is out of tolerance. (Low) | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:29 | Biopsy Guidance | Issued when the 12V bus voltage is out of tolerance. (High) | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:30 | Biopsy Guidance | Issued when the 12V bus voltage is out of tolerance. (Low) | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:31 | Biopsy Guidance | Issued when the Vcc bus voltage is out of tolerance. (High) | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:32 | Biopsy Guidance | Issued when the Vcc bus voltage is out of tolerance. (Low) | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:33 | Biopsy Guidance | Issued when invalid arguments are passed to the transmit message API. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:34 | Biopsy Guidance | Issued when invalid arguments are passed to the transmit extended announcement API. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:35 | Biopsy Guidance | Issued when the BCM CAN bus has been disabled due to multiple hardware errors. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:36 | Biopsy Guidance | Issued when the Gantry CAN bus has been disabled due to multiple hardware errors. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:37 | Biopsy Guidance | Issued when a buffer overrun has occurred while routing messages onto the BCM CAN bus. Data was lost. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:38 | Biopsy Guidance | Issued when a buffer overrun has occurred when trying to route messages onto the CANbus. Data was lost. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:39 | Biopsy Guidance | Issued when a buffer overrun has occurred when routing messages to the BGM Application Software. Data was lost. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:40 | Biopsy Guidance | Issued when overrun errors have been detected on the BCM CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:41 | Biopsy Guidance | Issued when overrun errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:42 | Biopsy Guidance | Issued when transmit errors have been detected on the BCM CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:43 | Biopsy Guidance | Issued when transmit errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:44 | Biopsy Guidance | Issued when receive errors have been detected on the BCM CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:45 | Biopsy Guidance | Issued when receive errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:46 | Biopsy Guidance | Issued when some transmit errors have been detected on the BCM CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:47 | Biopsy Guidance | Issued when some transmit errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:48 | Biopsy Guidance | Issued when receive errors have been detected on the BCM CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:49 | Biopsy Guidance | Issued when some receive errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:50 | Biopsy Guidance | Issued when the BGM fails POST. | Address the failure indicated in the POST result (i.e. voltage out of spec , etc.) If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:51 | Biopsy Guidance | Issued when the bootloader has been started because Application failed to load properly. | Possible bad code download, reload application code. If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:52 | Biopsy Guidance | Issued when the POST calibration data check fails (corrupted). | Recalibrate or reload stored configuration data. If the problem persists, collect all logs and forward to service as board replacement may be necessary.
34:53 | Biopsy Guidance | Issued when a voltage measures out of tolerance on the BGM. | Check all line voltages feeding the board. Inspect connections and on-board regulators for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
34:54 | Biopsy Guidance | Issued when a request to write data to the EEPROM has failed. | If the problem persists, collect all logs and forward to service as board replacement may be necessary.
34:55 | Biopsy Guidance | Issued when a request to read data from the EEPROM has failed. | If the problem persists, collect all logs and forward to service as board replacement may be necessary.
34:56 | Biopsy Guidance | Issued when invalid arguments are passed to the fault library. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:57 | Biopsy Guidance | Issued when an unexpected or unhandled interrupt has occurred. | Check all connections and board voltages. Collect all logs and forward to service as soon as possible. Board replacement may be necessary.
34:58 | Biopsy Guidance | Issued when a task stack overrun has been detected. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:59 | Biopsy Guidance | Issued when an unhandled or unexpected OS failure has been detected. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
34:60 | Biopsy Guidance | Issued when BGM software gets into a state it should not be in. | Check all connections and board voltages. Collect all logs and forward to service as soon as possible.
34:61 | Biopsy Guidance | Issued when the biopsy device is violating one or more safety margin limits. | Move the biopsy device outside of the safety margin limit. Recalibration may be required or needle parameters may be inaccurate if visual inspection would indicate non-infringement. If the problem persists, collect all logs and forward to service.
34:62 | Biopsy Guidance | Issued when a hardware limit reached and motion cannot continue. | Check calibrations and pot alignment. Recalibration may be required. If problem persists, collect all logs and forward to service.
34:63 | Biopsy Guidance | Issued when the BGM serial number does not match serial number in the STX matrix provided by the AWS. | Verify STX calibration, redo STX calibrations as necessary. If the problem persists, collect all logs and forward to service.
34:64 | Biopsy Guidance | Issued when one or more of the biopsy needle parameters are out of range. | Check for and correct all invalid needle parameters on the AWS. If the problem persists, collect all logs and forward to service.
34:65 | Biopsy Guidance | Issued when the BGM does not receive an ACCEPT or DENY to a command within the timeout period. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:66 | Biopsy Guidance | Issued when an invalid lateral mode setting was detected. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service. | bad data value
34:67 | Biopsy Guidance | Issued when a request to validate a target was made before the BGM was ready for validation. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:68 | Biopsy Guidance | Issued when lateral motion is requested without a Lat X value being set first. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:69 | Biopsy Guidance | Issued when lateral motion is requested without a Lat-X value being set first. | Check for compatible software versions. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
34:70 | Biopsy Guidance | Issued when motion in the wrong direction is detected. | Check the POT and motor connections for any defects or shorts. If the problem persists, board or part replacement may be required. | correct dir to move, 0 - left/up (adc lowers), 1 - right/down (adc grows) | mot | uint8_t | 0 - x, 1- y
36:1 | Switch | Issued when manual switch commands have not been received for over 100ms and motion cannot be maintained. | Inspect the cable assemblies for proper CAN connections and terminations. Inspect the log for missing, late, invalid, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
37:1 | C-arm Display | Issued when an unexpected or unhandled message is received from another module in the system. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If the problem persists, collect all logs and forward to service.
38:1 | Power Control | (obs)Issued when an unexpected and unhandled message is received at the msg router | No longer used.
38:2 | Power Control | (obs)Issued when a message is lost due to excessive CAN msg traffic | No longer used.
38:3 | Power Control | (obs)Issued when a message is lost due to excessive CAN msg traffic | No longer used.
38:4 | Power Control | (obs)Issued when a message is lost due to excessive CAN msg traffic | No longer used.
38:5 | Power Control | (obs)Issued when the Gantry CAN Tx task has had to wait for an empty Tx register | No longer used.
38:6 | Power Control | (obs)Issued when the AWS CAN Tx task has had to wait for an empty Tx register | No longer used.
38:7 | Power Control | (obs)Issued when the SVC CAN Tx task has had to wait for an empty Tx register | No longer used.
38:8 | Power Control | (obs)Issued when the task attempted to send a message with the port disabled | No longer used.
38:9 | Power Control | (obs)Issued when the task attempted to send a message with the port disabled | No longer used.
38:10 | Power Control | (obs)Issued when the task attempted to send a message with the port disabled | No longer used.
38:11 | Power Control | Issued when an error in one or more of the local voltage values has been detected. | Check the DC power circuits for defects or shorts.
38:12 | Power Control | Issued when an error in one or more of the power monitor values occurs. | Check AC power circuits for defects or shorts.
38:13 | Power Control | Issued when an On/Off Stuck Switch condition is detected during power up. | Inspect the mechanical and electrical components for defects or shorts.
38:14 | Power Control | Issued when a display port function has returned an error at the PMC. | Check angle displays, wiring, and connections for defects or shorts.
38:15 | Power Control | Issued when the PMC detects a fan failure (over current). | Check line voltages and on-board regulator. Inspect the fan and all connections. Replace the fan if necessary.
38:16 | Power Control | Issued when the E-Stop is tripped. | Check the E-Stop switches and VTA/rotation limit switches. Check E-Stop wiring.
38:17 | Power Control | Issued when the trip coil circuit is in a fault state. | Check the trip coil circuit for defects or shorts.
38:18 | Power Control | Issued when a fault has been detected in the emergency interlock circuit, as there is an indication of a set switch but no indication of which one is set. | Check for wiring defects or shorts, bad switches, and disconnected switches on E-Stops.
38:19 | Power Control | Issued when the thermal fuse has been tripped. | Check the thermal fused and circuit. Inspect the gantry for any mechanical or electrical problems.
38:20 | Power Control | (obs)Timeout waiting for NODE_SHUTDOWN_READY. Data[0] = node ID
38:21 | Power Control | Issued when the PMC has detected that another node has failed to respond to the shutdown prep message. | The PMC has detected that another node has rebooted or shutdown without responding to the shutdown prep message. Check the log for notifications or reset frames from other nodes. If the problem persists, collect all logs and forward to service.
38:22 | Power Control | Issued when power is detected after Power-Off. | Check for fused power relays on PMC board.
38:23 | Power Control | Issued 1 second before the circuit breaker trip is activated. | Inspect all power connections and wiring. Inspect all power control components for defects or shorts.
38:24 | Power Control | Issued when an emergency shutdown command has ben received from another node. | This is an indication that another node has issued a safety-related EMERGENCY-OFF. Inspect the logs for notifications from other nodes.
38:25 | Power Control | Issued when a CAN router fault is detected. **CLEAR will cause reboot** | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:26 | Power Control | Issued when an unexpected CAN_UP is detected. | A node has re-booted unexpectedly. Note: this is normal during code upgrades.
38:27 | Power Control | Issued when the VTA vertical travel or rotation E-Stop switch has been activated. | Check the VTA for proper operation. Check the VTA/rotation limit switches. Check E-Stop wiring.
38:28 | Power Control | Issued when the Left or Right Gantry E-Stop Button was pressed. | Check the Gantry E-Stop switches, reset as necessary. Check E-Stop wiring.
38:29 | Power Control | Issued when a fault was detected in the Left or Right Gantry E-Stop circuit. | Check E-stop switches. Check e-stop wiring.
38:30 | Power Control | Issued when the PMC cannot get valid data out of the Clock Chip. | Check of clock battery and replace if necessary. Possible bad clock chip on PMC board, replace board if the problem persists.
38:31 | Power Control | Issued when Power-timer control functions are enabled but the clock is not set | The PMC needs a proper clock sync from the AWS. Else, the PMC clock is bad.
38:32 | Power Control | Issued when an invalid time stamp is received from the AWS. | Fault will self-clear when a valid time stamp is received from the AWS. Check for a valid date/time setting on the AWS computer.
38:33 | Power Control | Issued when the detector (DET) is shut down because the DTC heartbeat is not sensed on the CAN bus. | Check for proper DTC operation. Check CAN connections/terminations. Check DTC, DET power supply circuits.
38:34 | Power Control | Issued when invalid arguments have been passed to the transmit message API. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
38:35 | Power Control | Issued when invalid arguments have been passed to transmit extended announcement API. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
38:36 | Power Control | Issued when a buffer overrun has occurred when trying to route messages onto the Gantry CAN bus. Data was lost. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
38:37 | Power Control | Issued when a buffer overrun has occurred when trying to route messages onto the AWS CAN bus. Data was lost. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
38:38 | Power Control | Issued when a buffer overrun has occurred when trying to route messages to the PMC application software. Data was lost. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
38:39 | Power Control | Issued when a bus overrun error has been detected on the Gantry CAN bus. | Inspect the cable assemblies on the Gantry for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:40 | Power Control | Issued when some transmit errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:41 | Power Control | Issued when transmit errors have been detected on the on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:42 | Power Control | Issued when the Gantry CAN bus has been disabled due to multiple hardware errors. | Inspect the cable assemblies on the Gantry for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:43 | Power Control | Issued when a bus overrun has been detected on the AWS CAN bus. | Inspect the cable assemblies between the PMC and the AWS for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:44 | Power Control | Issued when some transmit errors have been detected on the AWS CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:45 | Power Control | Issued when transmit errors have been detected on the AWS CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:46 | Power Control | Issued when the AWS CAN bus has been disabled due to multiple hardware errors. | Inspect the cable assemblies between the PMC and the AWS for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:47 | Power Control | Issued when some receive errors have been detected on the AWS CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:48 | Power Control | Issued when receive errors have been detected on the AWS CAN bus. | Inspect the cable assemblies between the PMC and the AWS for proper connections, breaks, shorts, and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:49 | Power Control | Issued when some receive errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:50 | Power Control | Issued when reveive errors have been detected on the Gantry CAN bus. | Inspect the cable assemblies for proper connections and CAN bus terminations. Inspect the log for missing, late, or failure responses from other modules. If problem persists, collect all logs and forward to service.
38:51 | Power Control | Issued when another node (i.e. GCB) sensed multiple switch closures and disabled this switch for safety. | Inspect the foot switch mechanical and electrical components for defects or shorts.
38:52 | Power Control | Issued when another node (i.e. GCB) sensed multiple switch closures and disabled this switch for safety. | Inspect the foot switch mechanical and electrical components for defects or shorts.
38:53 | Power Control | Issued when the left foot switch is stuck. (POST) | Inspect the foot switch mechanical and electrical components for defects or shorts.
38:54 | Power Control | Issued when the right foot switch is stuck. (POST) | Inspect the foot switch mechanical and electrical components for defects or shorts.
38:55 | Power Control | Issued when the left foot switch is stuck. (run-time fault) | Inspect the foot switch mechanical and electrical components for defects or shorts.
38:56 | Power Control | Issued when the right foot switch is stuck. (run-time fault) | Inspect the foot switch mechanical and electrical components for defects or shorts.
38:57 | Power Control | Spare | not used
38:58 | Power Control | Issued when the fan has been detected to be in a fault state for 1 minute or longer. | Check line voltages and on-board regulator. Inspect the fan and all connections. Replace the fan if necessary.
38:59 | Power Control | Issued when the PMC has detected another gantry node has reset. | The PMC has detected another node has reset, check the log for notifications or reset frames from other modules. If the problem persists, collect all logs and forward to service.
40:1 | Comp Dev Intfc | (obs)Issued when an unexpected and unhandled message is received at the msg router | Incompatible software versions.
40:2 | Comp Dev Intfc | Issued when the +5V or +15V line is out of tolerance | Check line voltages and on-board regulators. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
40:3 | Comp Dev Intfc | Issued when the PWB temperature is outside of the specified tolerance. | Check line voltages and on-board regulators. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
40:4 | Comp Dev Intfc | Issued when a request for paddle motion is denied due to the compression force. | Release compression. If no compression is applied and lockout is still in effect, recalibrate the compression force. If the problem persists, collect all logs and forward to service.
40:5 | Comp Dev Intfc | Issued when a stuck switch is detected. | Check for stuck switch. Inspect all housing components proper fit and connection. Check the board mechanical and electrical components for defects or shorts. If problem persists, collect all logs and forward to service.
40:6 | Comp Dev Intfc | Issued when a paddle shift is commanded and the paddle shift motor fails to reach the commanded position before the motion timeout is passed | Check for something physically blocking the shift mechanism. Check the integrity of the paddle shift mechanism. If the gears move freely and the motor can be driven, repeat the paddle shift calibration.
40:7 | Comp Dev Intfc | Issued when noise is detected on the compression thickness potentiometer ADC channel. | Check the compression thickness potentiometer connections and the potentiometer itself for defects or shorts. If the problem persists, replace the compression thickness potentiometer.
58:1 | Status | Issued when a board has failed POST. | Address the failure indicated in the POST result (i.e. voltage out of spec , etc.) If the problem persists, collect all logs and forward to service. Board replacement may be necessary. | POST result
58:2 | Status | Issued when the bootloader was started because the application failed to launch. | Possible bad code or failed application download, reload the application code. If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:3 | Status | Issued when a calibration data check fails (corrupted) during POST. | Load a saved configuration or restore the board defaults. Recalibration may be required. If the problem persists, collect all logs and forward to service.
58:4 | Status | Issued when a voltage has been measured to be out of tolerance. | Check line voltages and on-board regulators. Check the electrical components for defects or shorts. If problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:5 | Status | Issued when a request to write data to EEPROM has failed. | If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:6 | Status | Issued when a request to read data from EEPROM has failed. | If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:7 | Status | Issued when invalid arguments have been passed to the fault library. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:8 | Status | Issued when the hardware has triggered an interrupt that the software was not expecting. | Inspect the board and all connections for any mechanical or electrical defects or shorts. Check the board for proper operation. Collect all logs and forward to service. Board replacement may be necessary.
58:9 | Status | Issued when a stack overrun has been detected. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:10 | Status | Issued when an unexpected OS failure has been detected. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:11 | Status | Issued when the software gets into a state it should not be in. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:12 | Status | Issued when an error has been detected in the CAN driver software interface. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:13 | Status | Issued when an error has occurred while attempting to erase or write to a block of flash memory. | Re-attempt the software installation or configuration update. If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:14 | Status | Issued when manual switch commands have not been received in over 100ms. Motion cannot be maintained. | Inspect the cable assemblies for proper CAN connections and terminations. Inspect the log for missing, late, invalid, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
58:15 | Status | Issued when the hardware has detected a data overrun on the CAN bus. | Inspect the cable assemblies for proper CAN connections and terminations. Inspect the log for missing, late, invalid, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
58:16 | Status | Issued when the hardware has disabled the CAN bus due to multiple errors. | Inspect the cable assemblies for proper CAN connections and terminations. Inspect the log for missing, late, invalid, or failure responses from other modules in the system. If problem persists, collect all logs and forward to service.
58:17 | Status | Issued when invalid arguments have been passed to the transmit message API. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:18 | Status | Issued when a request to write data to NVM has failed. | If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:19 | Status | Issued when a request to read data from NVM has failed. | If the problem persists, collect all logs and forward to service. Board replacement may be necessary.
58:20 | Status | Issued when a software task has been halted. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:21 | Status | Issued when the foundation library has failed to initialize. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:22 | Status | Issued when the application software has failed to initialize. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:23 | Status | Issued when a corrupted configuration block is detected. | Reload a saved configuration file if available for the specific software version installed on the board or restore the board defaults. If the problem persists, collect all logs and forward to service.
58:24 | Status | Issued when a I2C-connected hardware peripheral is unable to communicate with processor. | Collect all logs and forward to service as soon as possible. No user or FE action possible.
58:25 | Status | Issued when a NVM initialization has failed. | If the problem persists, collect all logs and forward to service. Reload a saved configuration or restore the board defaults. Board replacement may be necessary.
58:26 | Status | Issued when a software task was unable to set an internal event. | Collect all logs and forward to service as soon as possible. No user or FE action possible.